Update OBOD-IDS probes from 1.04 to 1.05

Symptom:  OBOD-IDS probes can lose all communication with known good MultiLab meters.  Once communication is lost, the communication modules cannot be restored / recovered.  The probe will not operate with the meter.

Solution:  Update the functional OBOD-IDS communication module firmware from 1.04 to 1.05 before failure occurs.

The following four steps only need to be performed at the first update:

- Install the OBOD-IDS Update software onto a PC (Update provided by WTW).

- Install the MultiLab VCP driver from the YSI website:

https://www.ysi.com/customer-support/software-firmware-downloads/software

- Create a desktop shortcut to the OBOD IDS Update program installed above.

- Verify the 4010-1W used for this upgrade is at Software version 2.13 or newer; Press and hold Enter until the meter configuration menu is accessed.  Select System.  Select Service Information.  The software version will be displayed at the bottom of this window.  If the Software version is < 2.13, update the 4010-1W accordingly following the User manual Section 16.

Connect the devices

Connect the OBOD-IDS probe to a known good MultiLab 4010-1W meter using the IDS communication port.  Note the stir motor does not have to be connected.

Connect the 4010-1W / probe system to the PC using the USB communication cable.

Turn the system ON.

Perform the OBOD-IDS Update

Open the OBOD-IDS Update program.  The screen will appear as below.  Note the PC will automatically establish the correct Com port (Com3 in this example).

Click OK to begin the update.  The update will take approximately 1 minute.

Continued on next page

When the update is finished, the screen will appear as below.  Select File -> Exit.

Verify Update

Cycle power on the meter.  At the main screen press the F1 / “i” button twice to read the probe information.  The probe software (1.05) will be displayed at the bottom of this probe information window.